## 1. FOGOffset&MSG
from ME TPI reports, pick out the stability failed HI, retest the TPI.
then, keep the HI no moved, do the DFS
get the MicStabilityFOGOffset to be reference
combine it with MSG
| Name | FinalFrontMicStabilityFOGOffset
| 1800801007 | 1800801005 | 1800801010
| 4 | 4 | 4
| Name | FinalBackMicStabilityFOGOffset
| 1800801007 | 1800801005 | 1800801010
| 4 | 4 | 4
| Name | MicStabEXPOffAGCICFG
| 1800801007 | 1800801005 | 1800801010
| 1 | 1 | 1
VA877-DW stability failure
| Name | FinalFrontMicStabilityFOGOffset |  | and the unstable frequency can match the MSG_DFS
| 1800801240 | 1800801239 |  |  | #240: 3981 or 5012
| 1 | 0 |  |  | #239: 3981 or 5012
| Name | FinalBackMicStabilityFOGOffset
| 1800801240 | 1800801239
| 4 | 4
| Name | MicStabEXPOffAGCICFG
| 1800801240 | 1800801239
| 1 | 1
change the jig to Delhi, the result is the same
so jig is not the root cause
do the MicStabilityFOGOffset and MSG_DFS have connection?
[6 empty] | Name | FinalFrontMicStabilityFOGOffset
[6 empty] | 1800801007 | 1800801005 | 1800801010
[6 empty] | 4 | 4 | 4
[6 empty] | 1800801240 | 1800801239
[6 empty] | 1 | 0
[6 empty] | Name | FinalBackMicStabilityFOGOffset
[6 empty] | 1800801007 | 1800801005 | 1800801010
[6 empty] | 4 | 4 | 4
[6 empty] | 1800801240 | 1800801239
[6 empty] | 4 | 4
it seems their offset are same, but the MSG_DFS is different.
so this offset is for a reference when HI is stable
LS977-DW stability failure
[9 empty] | Name | FinalFrontMicStabilityFOGOffset
[9 empty] | 1458 | 2564 | 3568 | 0037
[9 empty] | 3 | 0 | 2 | 4
[9 empty] | Name | FinalBackMicStabilityFOGOffset
[9 empty] | 1458 | 2564 | 3568 | 0037
[9 empty] | 4 | 2 | 4 | 4
[9 empty] | Name | MicStabEXPOffAGCICFG
[9 empty] | 1458 | 2564 | 3568 | 0037
[9 empty] | 1 | 1 | 1 | 1
using the Delhi Jig
check if it can improve the performance
[9 empty] | Name | FinalFrontMicStabilityFOGOffset
[9 empty] | 3568 | Delhi Jig
[9 empty] | 2 | 4
[9 empty] | Name | FinalBackMicStabilityFOGOffset
[9 empty] | 3568 | Delhi Jig
[9 empty] | 4 | 4
[9 empty] | Name | FinalFrontMicStabilityFOGOffset
[9 empty] | 2564 | Delhi Jig
[9 empty] | 0 | 0
[9 empty] | Name | FinalBackMicStabilityFOGOffset
[9 empty] | 2564 | Delhi Jig
[9 empty] | 2 | 2
prove it again: jig is not the root cause

## 2. RS_vs_BT skin
compare the stability of Resound and beltone skin
FinalFrontMicStabilityFOGOffset
[5 empty] | BT | P5#458 | P6#240 | P6#239 | P6#237 | P6#238 | P6#236
[6 empty] | 1 | 1 | 0 | 2 | 4 | 4
[5 empty] | RS | P6#564 | P6#568 | P6#037
[6 empty] | 0 | 3 | 4
peak frequency:
[5 empty] | RS | 3k-4k Hz
[5 empty] | BT | 4k-5k Hz
FinalBackMicStabilityFOGOffset
[5 empty] | BT | P5#458 | P6#240 | P6#239 | P6#237 | P6#238 | P6#236
[6 empty] | 4 | 4 | 4 | 4 | 4 | 4
[5 empty] | RS | P6#564 | P6#568 | P6#037
[6 empty] | 1 | 4 | 4
replace failed BT HI skin with passed RS
P6 | FinalFrontMicStabilityFOGOffset |  |  | P6 | FinalBackMicStabilityFOGOffset
#037_original | 4 |  |  | #037_original | 4
#037_with BT skin | 3 |  |  | #037_with BT skin | 4
#037_recover | 4 |  |  | #037_recover | 4
#037_replace 2 | 0 |  |  | #037_replace 2 | 1
#037_replace 3 | 4 |  |  | #037_replace 3 | 2
#037_replace 4 | 0 |  |  | #037_replace 4 | 2
#037_recover_2 | 4 |  |  | #037_recover_2 | 4
first time, replace the skin with BT, the stability go worse a little bit
then recover to RS, stability return ok
replace to BT again, and repeat the disassembly, the performance not good
after several time skin replacement, once recover to RS skin, stability is good
For these two housing, RS is better ( RS:+1=1p)
replace failed BT HI skin with RS
P6 | FinalFrontMicStabilityFOGOffset |  |  | P6 | FinalBackMicStabilityFOGOffset
#240_original | 1 |  |  | #240_original | 4
#240_with RS skin | 4 |  |  | #240_with RS skin | 4
#240_recover | 4 |  |  | #240_recover | 4
#240_replace 2 | 4 |  |  | #240_replace 2 | 4
#240_replace 3 | 4 |  |  | #240_replace 3 | 4
#240_recover_2 | 4 |  |  | #240_recover_2 | 4
replace with this RS skin, stability become better obviously
recover to BT skin, stability problem seems fixed, the peak frequency shift
replace to RS several times, MSG keeps higher than BT
recover to BT again, the MSG lower
For these two housing, RS is better (RS:+1=2p)
replace the #240 failed HI with another RS skin
P6 | FinalFrontMicStabilityFOGOffset |  |  | P6 | FinalBackMicStabilityFOGOffset
#240_original | 1 |  |  | #240_original | 4
#240_with RS skin | 4 |  |  | #240_with RS skin | 4
#240_RS SkinHook | 4 |  |  | #240_RS SkinHook | 4
#240_recover2BT | 3 |  |  | #240_recover2BT | 4
HI #240 failed in front stability originally
replace with RS shell, it can pass the test
keep RS shell, replace the hook with the RS one, stability is good
recover to original shell and hook, front mic failed
test uses another RS skin, same conclusion, this BT skin is worse (RS:+1=3p)
P6 | FinalFrontMicStabilityFOGOffset |  |  | P6 | FinalBackMicStabilityFOGOffset
#564_original | 0 |  |  | #564_original | 1
#564_with BT skin | 0 |  |  | #564_with BT skin | 0
#564_BT SkinHook | 1 |  |  | #564_BT SkinHook | 0
#564_recover2RS | 0 |  |  | #564_recover2RS | 3
athough this HI failed the test
but replacing with BT skin makes stability of #564 worse
in a way, we can make the conclusion: #564 skin is better (RS:+1=4p)
trying to find what root cause makes #240 skin worse
#236 vs #240:
they can pass both, but #236 is better a lot
between #236 and #240, swap their component to figure out root cause
swap hook: no obviouse impact on stability performance
swap shell: MSG drop down a lot, shell of #240 is worse
swap shell and also pin: no big change from pin
root cause is the shell of #240