## Cover page
| Form 0301660 REV.A | TEMPL, TPI Validation Report
Validation Report for TPI (pre)Release:<ModelNames>
| 1 | Purpose
| 1.1 | The purpose of this document is to report the result of the software validation of <ModelNames> TPI
| 2 | Test Information
|  | Eng: | <EngineerName>
|  | Model Name: | <ModelNames>
|  | Date: | <Date>
|  | No. Of Units: | <TotalNo.OfUnits>
|  | Test Software Platform Ver: | <Version>
|  | Test Framework Ver: | <FrameworkVersion>
|  | Test Program Ver: | <TpiVersion>
|  | FDI Version: | <FdiVersion>
|  | Hybrid Version: | <HybridVersion>
|  | Test Analyzer: | <TestSystemName>
| 3 | References/Forms
| 3.1 | 0070290 V&V GUIDANCE PROCEDURE
| 3.2 | 0013050 Proc. Test software packages release (Fonix II, Audiomaster scripts)
| 3.3 | 7-50-2010 Audiomaster Custom User Manuel
| 3.4 | 7-26-6720 ATC 4000 Audio Test Chamber
| 4 | Equipment / fixturing / gauging
| 4.1 | List all equipment, fixturing and gauges used during the validation. This should include all calibration information.
|  | Identification # | Description | Last Cal. Date | Next Cal. Date
DSA6000
DSA6000
DSA6000
| 5 | Attachments to the reports
| 5.1 | <Aattachments>
| 6 | Summary&Conclusion
| 6.1 | All 30 devices are passed TPI test ,Framework version <FrwVersion> and Test program version <TpiVersion> are verified that can be used to test <ModelNames> devices .

## LS977-DW
Validation Report for TPI (pre)Release:LS977-DW
| Eng: | dendu
| Model Name: | LS977-DW
| Date: | 2017-12-01
| No. Of Units: | 50
| Test Software Platform Ver: | 2012 SP1,6.3.0
| Test Framework Ver: | 6.0.9
| Test Program Ver: | 1.1.7
| TPI Change Code: | B299-C2CC-2AEB-8998-1EE9
| FDI Version: | RS.P5.Top.BeBTE13DWNC.5.0
| Device Options: | VC DAI TC PB TBR WL MFI
| Hybrid Version: | C4.5 (34DB-45DA-6107)
| Test Type: | Final
LS977-DW
| Test Item |  | Spec |  |  | Test Data
| Prg | Description | Unit | Minimum | Maximum | Average | STD DEV | Min | Max | Cpl | Cpu | Cpk | 1231231230 | 1800800100 | 1800800973 | 1800800974 | 1800800975 | 1800800976 | 1800800977 | 1800800978 | 1800800979 | 1800800980 | 1800800981 | 1800800982 | 1800800983 | 1800800984 | 1800800985 | 1800800987 | 1800800988 | 1800800989 | 1800800990 | 1800800991 | 1800800992 | 1800800993 | 1800800994 | 1800800995 | 1800800996 | 1800800997 | 1800800998 | 1800800999 | 1800801000 | 1800801002 | 1800801003 | 1800801004 | 1800801005 | 1800801006 | 1800801007 | 1800801008 | 1800801009 | 1800801010 | 1800801011 | 1800801012 | 1800801013 | 1800801014 | 1800801015 | 1800801016 | 1800801017 | 1800801018 | 1800801019 | 1800801020 | 1800801021 | 1800801022 | Unit | Description
| P1 | Write Device Image | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | Write Device Image
| P1 | Input Calibration | dB | -5 | 8 | 1.1988 | 0.34692 | 0.48 | 1.82 | 5.956 | 6.5348 | 5.956 | 1.45 | 0.81 | 1.16 | 1.42 | 1.33 | 1.42 | 1.28 | 1.54 | 1.28 | 0.91 | 1.33 | 1.56 | 1.3 | 1.3 | 1.82 | 1.3 | 1.45 | 1.14 | 1.21 | 1.66 | 1.54 | 1.23 | 0.67 | 1.12 | 1.73 | 0.84 | 0.88 | 1.35 | 0.81 | 1.56 | 0.79 | 1.3 | 1.61 | 0.67 | 1.45 | 0.93 | 0.84 | 1.73 | 1.42 | 1.52 | 0.84 | 0.48 | 1.02 | 0.58 | 0.81 | 0.93 | 0.88 | 1.59 | 1.52 | 0.63 | dB | Input Calibration
| P1 | Output Calibration | dB | -4 | 9 | 1.89 | 0.41682 | 0.88 | 3.18 | 4.7102 | 5.6858 | 4.7102 | 1.01 | 2.1 | 1.75 | 1.63 | 2.06 | 1.44 | 1.34 | 2.01 | 1.74 | 1.95 | 2.11 | 1.96 | 1.69 | 2.2 | 2.46 | 1.74 | 1.92 | 1.74 | 1.45 | 1.89 | 2.3 | 2.12 | 1.74 | 1.85 | 2.15 | 1.85 | 1.46 | 2.83 | 2.02 | 2.59 | 2.37 | 1.99 | 1.99 | 2.01 | 1.89 | 2.13 | 1.65 | 3.18 | 1.45 | 2.01 | 1.95 | 1.71 | 1.39 | 0.88 | 2.04 | 1.77 | 1.92 | 2.04 | 1.02 | 2.01 | dB | Output Calibration
| P1 | Microphone Matching | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | Microphone Matching
| P1 | Wireless Connection | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | Wireless Connection
| P2 | OSPL90 Max | dBSPL | 119.39 | 127.39 | 124.038 | 0.36932 | 123.3 | 124.68 | 4.1951 | 3.0254 | 3.0254 | 124.27 | 123.38 | 124.11 | 124.39 | 123.95 | 124.39 | 124.43 | 124.43 | 124.3 | 123.9 | 124.22 | 124.26 | 124.35 | 124.07 | 124.26 | 124.29 | 124.45 | 124.19 | 124.22 | 124.65 | 124.08 | 123.99 | 123.77 | 124.29 | 124.16 | 123.53 | 124.02 | 123.52 | 123.33 | 123.8 | 123.73 | 124.37 | 124.31 | 123.7 | 124.49 | 123.85 | 124.15 | 123.8 | 124.52 | 124.44 | 123.53 | 123.3 | 123.69 | 123.98 | 123.41 | 123.61 | 123.62 | 124.14 | 124.68 | 123.58 | dBSPL | OSPL90 Max
| P2 | OSPL90 HFA | dBSPL | 116.77 | 124.77 | 119.633 | 0.43107 | 118.67 | 120.78 | 2.2139 | 3.9723 | 2.2139 | 120.78 | 119.08 | 119.75 | 120.05 | 119.55 | 120.09 | 120.21 | 119.71 | 119.85 | 119.36 | 119.59 | 119.87 | 119.91 | 119.34 | 119.72 | 119.81 | 119.86 | 119.56 | 119.89 | 120.08 | 119.54 | 119.34 | 119.15 | 119.51 | 119.8 | 119.4 | 119.96 | 118.96 | 119.17 | 119.43 | 118.67 | 119.73 | 119.79 | 118.93 | 119.94 | 119.27 | 119.68 | 119.11 | 120.28 | 119.64 | 119.2 | 119.21 | 120.05 | 120.01 | 119.1 | 119.6 | 119.37 | 119.69 | 120.76 | 119.3 | dBSPL | OSPL90 HFA
| P2 | OSPL90 500Hz | dBSPL | 112.74 | 122.74 | 116.5954 | 0.69178 | 114.95 | 117.57 | 1.8577 | 2.9608 | 1.8577 | 117.52 | 115.9 | 117.13 | 117.37 | 116.73 | 117.3 | 117.33 | 117.23 | 117.1 | 115.89 | 116.78 | 117.2 | 117.1 | 116.54 | 116.88 | 116.97 | 117.14 | 116.77 | 117.27 | 117.46 | 116.86 | 116.88 | 115.94 | 116.82 | 117.19 | 115.96 | 116.29 | 115.09 | 115.93 | 115.97 | 114.95 | 117.08 | 117.23 | 116.03 | 117.13 | 115.96 | 116.28 | 115.6 | 117.34 | 117.07 | 116.02 | 115.61 | 116.38 | 116.62 | 115.87 | 116.2 | 116 | 117.16 | 117.57 | 115.13 | dBSPL | OSPL90 500Hz
| P2 | OSPL90 1600Hz | dBSPL | 113.94 | 123.94 | 117.8506 | 0.41132 | 117.05 | 119.15 | 3.1691 | 4.9348 | 3.1691 | 119.15 | 117.53 | 117.88 | 118.14 | 117.71 | 118.2 | 118.4 | 117.66 | 117.88 | 117.94 | 117.54 | 117.9 | 117.98 | 117.34 | 117.89 | 118.03 | 117.95 | 117.8 | 118.09 | 118.03 | 117.49 | 117.37 | 117.45 | 117.66 | 117.93 | 117.74 | 118.47 | 117.54 | 117.59 | 117.83 | 117.3 | 117.62 | 117.84 | 117.05 | 118.12 | 117.49 | 117.95 | 117.34 | 118.38 | 117.75 | 117.44 | 117.76 | 118.6 | 118.27 | 117.43 | 117.97 | 117.63 | 117.77 | 119.01 | 117.7 | dBSPL | OSPL90 1600Hz
| P2 | 50dB Template | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | 50dB Template
| P2 | FOG50 HFA Gain | dB | 47.9 | 57.9 | 52.526 | 0.30815 | 51.85 | 53.1 | 5.0041 | 5.8132 | 5.0041 | 52.46 | 52.7 | 52.61 | 52.76 | 52.61 | 52.39 | 52.27 | 52.35 | 52.38 | 52.81 | 52.53 | 52.26 | 52.44 | 52.56 | 52.31 | 52.11 | 52.5 | 52.36 | 52.43 | 52.76 | 51.99 | 52.35 | 51.98 | 52.37 | 52.01 | 52.9 | 52.37 | 53.1 | 52.85 | 53.02 | 53.04 | 52.59 | 52.55 | 52.4 | 52.52 | 52.62 | 52.73 | 52.97 | 52.75 | 52.14 | 52.07 | 52.85 | 52.64 | 52.98 | 52.27 | 52.84 | 52.79 | 51.85 | 52.24 | 52.92 | dB | FOG50 HFA Gain
| P2 | FOG50 1600Hz Gain | dB | 45.9 | 55.9 | 50.5536 | 0.35288 | 49.86 | 51.27 | 4.3958 | 5.0502 | 4.3958 | 50.44 | 51.22 | 50.4 | 50.25 | 50.51 | 50.28 | 50.37 | 50.15 | 50.2 | 51.02 | 50.45 | 50.36 | 50.36 | 50.44 | 50.29 | 50.39 | 50.37 | 50.33 | 50.32 | 50.37 | 50.31 | 50.28 | 49.86 | 50.4 | 50.15 | 51.01 | 51.24 | 51.15 | 50.95 | 51.09 | 51.27 | 50.33 | 50.41 | 50.58 | 50.63 | 50.88 | 50.79 | 50.95 | 50.3 | 50.23 | 50.58 | 51.05 | 50.93 | 50.55 | 50.62 | 50.49 | 50.59 | 49.9 | 50.49 | 51.15 | dB | FOG50 1600Hz Gain
| P2 | FOG50 Max Gain | dB | 52.29 | 62.29 | 58.1492 | 0.87361 | 56.5 | 61.54 | 2.2356 | 1.58 | 1.58 | 56.83 | 57.45 | 57.96 | 57.87 | 58.12 | 57.73 | 57.71 | 57.72 | 57.84 | 61.54 | 58.48 | 57.65 | 58.1 | 59.48 | 56.5 | 57.81 | 58.03 | 58.41 | 57.59 | 59.34 | 57.9 | 58.33 | 58.64 | 57.72 | 57.55 | 57.98 | 56.9 | 59.54 | 57.81 | 58.65 | 60.3 | 58.02 | 58.36 | 57.88 | 59.52 | 58.32 | 58.24 | 58.7 | 58.26 | 57.6 | 57.52 | 57.32 | 57.32 | 58.04 | 58.01 | 58.12 | 58.51 | 57.61 | 57.37 | 59.26 | dB | FOG50 Max Gain
| P2 | Telecoil Stability | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | Telecoil Stability
| P3 | HFA 60dB | dBSPL | 98.04 | 108.04 | 102.6826 | 0.21111 | 102.25 | 103.15 | 7.3305 | 8.4592 | 7.3305 | 102.49 | 102.84 | 102.65 | 102.67 | 102.62 | 102.55 | 102.5 | 102.57 | 102.5 | 102.85 | 102.77 | 102.55 | 102.59 | 102.66 | 102.56 | 102.48 | 102.68 | 102.48 | 102.41 | 102.69 | 102.49 | 102.53 | 102.3 | 102.52 | 102.35 | 102.93 | 102.91 | 102.99 | 102.87 | 102.95 | 102.83 | 102.74 | 102.58 | 102.73 | 102.65 | 102.91 | 102.98 | 103.12 | 102.66 | 102.36 | 102.67 | 102.83 | 102.82 | 102.92 | 102.72 | 102.85 | 102.95 | 102.25 | 102.46 | 103.15 | dBSPL | HFA 60dB
| P3 | 60dB Template | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | 60dB Template
| P3 | Mic2 Offset | dB | -2 | 2 | 0.1194 | 0.10411 | -0.54 | 0.24 | 6.7859 | 6.0213 | 6.0213 | 0.16 | 0.22 | 0.04 | 0.15 | 0.13 | 0.09 | 0.15 | 0.14 | 0.12 | 0.14 | 0.17 | 0.11 | 0.13 | 0.15 | 0.08 | 0.14 | 0.14 | 0.17 | 0.13 | 0.17 | 0.1 | 0.07 | -0.54 | 0.09 | 0.11 | 0.16 | 0.24 | 0.2 | 0.09 | 0.17 | 0.19 | 0.15 | 0.18 | 0.09 | 0.18 | 0.09 | 0.13 | 0.05 | 0.17 | 0.11 | 0.09 | 0.18 | 0.16 | 0.13 | 0.08 | 0.07 | 0.13 | 0.11 | 0.13 | 0.13 | dB | Mic2 Offset
| P3 | Fmin/F1 | Hz |  |  | 100 | 0 | 100 | 100 |  |  |  | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | Hz | Fmin/F1
| P3 | Fmax/F2 | Hz |  |  | 6476.2254 | 97.3786 | 6334.8 | 6698.38 |  |  |  | 6334.8 | 6680.72 | 6404.07 | 6392.65 | 6432.4 | 6404.93 | 6415.88 | 6394.6 | 6373.01 | 6674.48 | 6404.05 | 6394.44 | 6397.27 | 6402.21 | 6443.26 | 6422.05 | 6419.58 | 6434.74 | 6445.16 | 6416.18 | 6392.05 | 6415.62 | 6639.89 | 6414 | 6433.81 | 6496.19 | 6588.75 | 6676.93 | 6622.84 | 6546.13 | 6698.38 | 6372.72 | 6423.07 | 6482.92 | 6433.04 | 6613.23 | 6600.48 | 6579.66 | 6354.22 | 6429.69 | 6496.71 | 6629.72 | 6533.71 | 6453.35 | 6526.59 | 6464.62 | 6459.74 | 6424.53 | 6424.66 | 6497.54 | Hz | Fmax/F2
| P3 | % THD at 500Hz | % | 0 | 3.36 | 0.5368 | 0.19202 | 0.29 | 1.36 | 0.93185 | 4.9009 | 0.93185 | 0.44 | 0.61 | 0.82 | 0.45 | 0.43 | 0.47 | 0.46 | 0.4 | 0.38 | 0.57 | 0.64 | 0.44 | 0.51 | 0.5 | 0.52 | 0.5 | 0.36 | 0.34 | 0.42 | 0.64 | 0.39 | 0.48 | 0.55 | 0.43 | 0.55 | 0.5 | 0.57 | 1.19 | 0.53 | 0.45 | 1.36 | 0.46 | 0.56 | 0.29 | 0.77 | 0.73 | 0.41 | 0.34 | 0.45 | 0.4 | 0.48 | 0.73 | 0.47 | 0.57 | 0.48 | 0.5 | 0.63 | 0.56 | 0.37 | 0.74 | % | % THD at 500Hz
| P3 | % THD at 800Hz | % | 0 | 3.84 | 0.6072 | 0.10795 | 0.37 | 0.93 | 1.8749 | 9.9824 | 1.8749 | 0.6 | 0.73 | 0.68 | 0.59 | 0.63 | 0.62 | 0.53 | 0.53 | 0.63 | 0.52 | 0.76 | 0.57 | 0.6 | 0.57 | 0.43 | 0.46 | 0.55 | 0.46 | 0.59 | 0.73 | 0.64 | 0.55 | 0.37 | 0.54 | 0.44 | 0.59 | 0.69 | 0.86 | 0.71 | 0.61 | 0.93 | 0.51 | 0.58 | 0.58 | 0.65 | 0.68 | 0.73 | 0.52 | 0.64 | 0.5 | 0.62 | 0.68 | 0.66 | 0.69 | 0.57 | 0.67 | 0.63 | 0.47 | 0.5 | 0.77 | % | % THD at 800Hz
| P3 | % THD at 1600Hz | % | 0 | 3.65 | 0.5058 | 0.072621 | 0.35 | 0.64 | 2.3216 | 14.432 | 2.3216 | 0.62 | 0.52 | 0.56 | 0.57 | 0.51 | 0.43 | 0.54 | 0.48 | 0.59 | 0.35 | 0.54 | 0.58 | 0.59 | 0.6 | 0.46 | 0.5 | 0.46 | 0.49 | 0.48 | 0.47 | 0.51 | 0.5 | 0.47 | 0.57 | 0.44 | 0.49 | 0.43 | 0.39 | 0.37 | 0.64 | 0.42 | 0.54 | 0.41 | 0.52 | 0.61 | 0.46 | 0.49 | 0.64 | 0.5 | 0.47 | 0.47 | 0.38 | 0.49 | 0.59 | 0.54 | 0.55 | 0.57 | 0.6 | 0.5 | 0.39 | % | % THD at 1600Hz
| P3 | EIN | dBSPL | 0 | 23.45 | 23.0586 | 0.37158 | 21.99 | 23.42 | 20.685 | 0.35111 | 0.35111 | 23.18 | 23.3 | 23.11 | 22.9 | 23.11 | 22.73 | 22.72 | 23.2 | 23.06 | 23.39 | 23.14 | 23.23 | 22.94 | 23.35 | 23.32 | 23.14 | 23.34 | 23.04 | 23.42 | 23.2 | 23.4 | 23.42 | 23.4 | 23.4 | 23.36 | 23.29 | 23.38 | 23.12 | 23.34 | 23.18 | 23.42 | 22.59 | 23.38 | 22.85 | 23.05 | 23.29 | 22.95 | 23.34 | 22.81 | 23.42 | 22.91 | 22.2 | 22.35 | 21.99 | 22.87 | 22.65 | 22.05 | 23.39 | 22.93 | 22.38 | dBSPL | EIN
| P3 | EIN (Back Mic) | dBSPL | 0 | 23.45 | 21.63 | 0.60689 | 20.54 | 23.38 | 11.8803 | 0.99963 | 0.99963 | 20.54 | 21.6 | 21.63 | 21.68 | 21.8 | 20.73 | 20.92 | 21.23 | 21.36 | 22.27 | 21.13 | 21.09 | 21.42 | 22.03 | 22.26 | 21.94 | 21.67 | 21.77 | 21.52 | 20.75 | 23.38 | 21.46 | 21.84 | 21.78 | 20.95 | 22.3 | 20.69 | 21.83 | 21.82 | 20.97 | 22.42 | 20.67 | 22.88 | 20.93 | 21.44 | 22.97 | 21.77 | 22.13 | 21.07 | 22.11 | 21.18 | 21.54 | 21.45 | 21.38 | 22.01 | 21.95 | 21.5 | 22.32 | 21.54 | 21.88 | dBSPL | EIN (Back Mic)
| P3 | Current Drain | mA | 0.5 | 1.44 | 1.2216 | 0.030194 | 1.19 | 1.3 | 7.9663 | 2.4111 | 2.4111 | 1.2 | 1.21 | 1.22 | 1.2 | 1.21 | 1.2 | 1.19 | 1.2 | 1.27 | 1.22 | 1.28 | 1.27 | 1.21 | 1.21 | 1.27 | 1.2 | 1.2 | 1.2 | 1.2 | 1.21 | 1.19 | 1.21 | 1.26 | 1.21 | 1.21 | 1.22 | 1.19 | 1.26 | 1.28 | 1.22 | 1.25 | 1.22 | 1.22 | 1.2 | 1.21 | 1.3 | 1.2 | 1.22 | 1.2 | 1.19 | 1.29 | 1.22 | 1.2 | 1.2 | 1.21 | 1.21 | 1.2 | 1.2 | 1.24 | 1.28 | mA | Current Drain
| P3 | SPLIV HFA | dBSPL | 98.51 | 110.51 | 104.8754 | 0.73389 | 103.42 | 106.69 | 2.8912 | 2.5592 | 2.5592 | 104.17 | 105.56 | 104.21 | 104.26 | 104.34 | 104.43 | 104.37 | 104.57 | 104.23 | 105.44 | 104.62 | 104.64 | 104.45 | 104.42 | 104.66 | 104.21 | 104.44 | 104.09 | 104.03 | 104.66 | 103.42 | 104.13 | 105.34 | 104.08 | 104.51 | 105.64 | 105.75 | 106.15 | 105.51 | 106.4 | 105.48 | 104.48 | 104.52 | 105.38 | 104.38 | 105.55 | 105.73 | 106.69 | 104.28 | 103.86 | 105.44 | 105.38 | 105.54 | 105.26 | 105.48 | 105.55 | 105.66 | 104.2 | 104.4 | 105.78 | dBSPL | SPLIV HFA
| P3 | RTLS/ETLS | dB | -4.53 | 7.47 | 2.1924 | 0.57483 | 0.93 | 3.58 | 3.8982 | 3.0604 | 3.0604 | 1.67 | 2.72 | 1.56 | 1.59 | 1.72 | 1.88 | 1.87 | 1.99 | 1.73 | 2.58 | 1.85 | 2.09 | 1.85 | 1.76 | 2.1 | 1.73 | 1.77 | 1.61 | 1.62 | 1.97 | 0.93 | 1.6 | 3.04 | 1.56 | 2.16 | 2.72 | 2.84 | 3.15 | 2.63 | 3.44 | 2.66 | 1.74 | 1.93 | 2.65 | 1.73 | 2.64 | 2.75 | 3.58 | 1.62 | 1.51 | 2.77 | 2.55 | 2.72 | 2.35 | 2.76 | 2.7 | 2.71 | 1.95 | 1.94 | 2.63 | dB | RTLS/ETLS
| P4 | Front Mic Stability Status | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | Front Mic Stability Status
| P4 | Back Mic Stability Status | [10 empty] | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed | Passed |  | Back Mic Stability Status