## Slide 1: DFS -> MSG
Chandler
Jan, 2018

## Slide 2
Theory
Validate
Execute

## Slide 3: Impulse response
MLS, maximum length sequence, this is the sound you hear when you press calibrate.
  So, actually MLS is the stimulus of this feedback system which we can specify amplitude in ALT
DFSInit.init_front = accumulated response on front mic.
  It is the response of the feedback path in time domain. Like, it is the voltage output of mic during the calibration.
Accumulated response convolved with the MLS -> impulse response (compensated with some circular delay ect.)

## Slide 4
DFSInit.init_resp_front = impulse response on front mic
  Impulse response tells us the characteristics of the feedback path (transfer function from receiver to microphone)
  We usually deal with the feedback path in frequency domain, and since the impulse response is in the time domain. Then we make an FFT to obtain the frequency characteristics. -> msg_dfs_off (system cal. not included)

## Slide 5: Red line in ALT
To obtain MSG in ALT, we transform init_resp_front to frequency domain and included the system and production calibrations (or just take msg_dfs_off + calibration = msg)
In case of dual mic, it is the worse case between front and rear curves that determines msg for each frequency.

## Slide 6: Validate
Measure FOG50 & OPL, combine them -> MSG_opl = FOG - OPL
Measure the gain when setting is 0 -> MSG_ALT = msg + gain0
  This msg is not equal to the MSG_OPL, because H bridge is not included. Msg here is the maximum gain setting we can set in HI
Compare two MSG