## Slide 1: Investigate RHI issue in production
Logan & Chandler
2017.11.30

## Slide 2: The problems of RHI found right now
Deviation

## Slide 3: deviation = f( HI, chamber, jig+rec, coupler+rub)
jig+rec
coupler+rub

## Slide 4: deviation = f( HI, chamber, jig+rec, coupler+rub)
General test

## Slide 5: deviation = f( HI, chamber, jig+rec, coupler+rub)
HI:PB, either short or long, will shift the results in two status.
Make a requirement with ME to make sure this is under control.
Notes: PB will make the deviation, it can be improved by mechanic

## Slide 6: deviation = f( HI, chamber, jig+rec, coupler+rub)
Switch 3 chambers
deviation = f( HI, chamber, jig+rec, coupler+rub)
Notes: I have translated the jig each test in one same chamber, it seems the peak will shift once the jig is moved, even just translated

## Slide 7: deviation = f( HI, jig+rec, coupler+rub)
Notes: we think it is not just caused by the jig

## Slide 8: deviation = f( HI, jig+rec, coupler+rub)
reinsert the coupler can make deviation
J0+C1: **
J0+C2: ****
J0+C3: ***
J1+C1: *
J1+C2: *
J1+C3: **
Performance is better in jig0 whatever coupler
we can still find a good group, so maybe it can be avoid by a good jig and coupler.
Notes: Try different jig and coupler groups, the more star, the better performance of deviation

## Slide 9: deviation = f( HI, jig+rec, coupler+rub, re-insert)
Notes: Remember the deviation of this group j1 + c1
Reinsert 3 times, but carefully, the deviation will be smaller
Insert casually, or push the coupler down to touch the holder, or rotate the coupler, the deviation will be go bad.
But, no matter what you do , I can reinsert it hard to make it recover to be good.

## Slide 10: deviation = f( HI, jig+rec, coupler+rub, re-insert)
Insert the figure of jig design
Insert the figure of adaptor

## Slide 11: msg = f( HI, jig+rec, coupler+rub)